## All results
Pre-acceptance criteria not met
|  | Sensitivity ≥ 90% AND specificity ≥ 95% |  | NCP: | nucleocapsid protein
|  | Sensitivity < 90% OR specificity < 95% |  | RBD: | receptor binding domain |  | *: days post symptoms (PS)
|  | Sensitivity < 90% AND specificity < 95% |  | ND: | not determined in the IFU
Number | Manufacturer | Test name | Type | target | Date (end of study) | Belgian lab | <- IgG or IgG/IgM combination, compared to PCR -> | [5 empty] | <- IFU Data -> |  | CE mark? | Remarks
[7 empty] | # positive samples | Sensitivity t0 OR 0-7 days PS | Sensitivity t7 OR 8-14 days PS | Sensitivity t14 OR >14 days PS | # negative samples | Specificity | Highest sensitivity | Highest Specificity
1 | Zentech | quickZen COVID-19 IgM/IgG rapid test | rapid | RBD | 2020-04-07 | CHU Liège | 73 | 54.1 | - | 100 | 75 | 100 | 100 | 97 | √ | Overall sensitivity: | N | 73 | TP | % | Overall specificity: | N | 75 | TN | %
[5 empty] | 2020-04-23 | UCL & ULB | 124 | 34.48 | 77.4 | 90.9 | 72 | 100 |  |  |  |  | 197 | 112.716 | 185.716 | 94.2721 |  | 147 | 72 | 147 | 100
2 | Avioq | Novel Coronavirus (2019-nCoV) Antibody IgG/IgM Assay (Colloidal Gold) | rapid |  | ND | UCL & ULB | 117 | 27.7 | 53.5 | 93.3 | 79 | 94.9 | no IFU |  | √
3 | BMT Bio marketing | Lab on Time 2019-nCoV IgG/IgM rapid test cassette | rapid |  | 2020-04-20 | UCL & ULB | 124 | 37.9 | 75.8 | 93.93 | 72 | 100 | no IFU |  | ?
4 | Euroimmun | Anti-SARS-CoV-2 ELISA IgG | ELISA | S1 | 2020-04-23 | UCL & ULB | 124 | 17.2 | 66.12 | 90.9 | 72 | 98.6 | 100 | 98.6 | √ | Overall sensitivity: | N | 112.716 | TP | % | Overall specificity: | N | 70.992 | TN | %
[5 empty] | 2020-07-31 | KULeuven* | 223 | 18.6 | 60.2 | 88.1 | 113 | 96.5 |  |  |  |  | 713 | 196.463 | 632.173 | 88.6638 |  | 425 | 109.045 | 414.4 | 97.5059
[5 empty] | 2020-05-19 | KULeuven | 153 | 21.6 | 55.1 | 89.5 | 103 | 96.1 | [5 empty] | 136.935 |  |  |  |  | 98.983
[5 empty] | 2020-05-25 | HIS-IZZ | 44 | - | - | 96 | 81 | 98 | [5 empty] | 42.24 |  |  |  |  | 79.38
[5 empty] | 2020-05-15 | AZ Delta* | 169 | 41.5 | 88.1 | 85.1 | 56 | 100 | [5 empty] | 143.819 |  |  |  |  | 56
5 | Hangzhou Clongene Biotech | Clungene COVID -19 IgM/IgG Rapid Test Cassette | rapid | envelope | 2020-05-19 | KULeuven | 153 | 35.1 | 64.1 | 97.4 | 103 | 90.3 | no IFU |  | √
6 | Zhejiang Orient Gene Biotech | COVID-19 IgG/IgM Rapid Test Cassette (Whole Blood/Serum/Plasma) | rapid | ND | 2020-05-19 | KULeuven | 153 | 46 | 80.8 | 97.4 | 103 | 91.3 | 97.2 | 100 | √ | Overall sensitivity: | N | 149.022 | TP | % | Overall specificity: | N | 94.039 | TN | %
[5 empty] | 2020-05-15 | AZ Delta | 171 | 52.5 | 93.48 | 92.94 | 55 | 92.86 |  |  |  |  | 324 | 158.9274 | 307.9494 | 95.0461 |  | 158 | 51.073 | 145.112 | 91.843
7 | Multi-G | COVID-19 IgG/IgM Ab Test Cassette (Whole blood/Serum/Plasma) | rapid | NCP | 2020-06-17 | ZOL | [6 empty] | 100 | 98 | √ | OEM = HANGZHOU ALLTEST BIOTECH
[5 empty] | 2020-05-19 | KULeuven | 153 | 43.2 | 71.8 | 97.4 | 103 | 88.3
8 | PRIMA Lab | COVID-19 IgG/IgM Rapid Test | rapid | ND | 2020-05-19 | KULeuven | 153 | 56.8 | 79.5 | 100 | 103 | 85.4 | 100 | 98 | √
9 | Shenzhen YHLO Biotech | iFlash-SARS-CoV-2 IgG | CLIA | ND | 2020-04-29 | Eupen | 108 | 34.1 | 73.3 | 90.9 | 75 | 100 | 97.5 | 96.3 | √
10 | VivaChek | VivaDiag SARS-CoV-2 IgM/IgG Rapid Test | rapid |  | 2020-05-19 | KULeuven | 153 | 35.1 | 64.1 | 97.4 | 103 | 99 | no IFU |  | √
11 | Liming Bio | StrongStep SARS-CoV-2 lgM/lgG Antibody Rapid Test | rapid | ND | 2020-05-19 | KULeuven | 153 | 35.1 | 66.7 | 97.4 | 103 | 98.1 | 93.3 | 98.7 | √
12 | Dynamiker Biotechnology (Tianjin) | 2019 nCOV IgG/IgM Rapid Test | rapid | NCP | 2020-05-19 | KULeuven | 153 | 46 | 66.7 | 97.4 | 103 | 95.2 | no IFU |  | √
13 | Cypress Diagnostics | COVID-19 IgM/IgG Rapid Test | rapid | ND |  | AZ Klina & GZA | [6 empty] | 95.35 | 100 | √
14 | Roche | Elecsys Anti-SARS-CoV-2 | ECLIA | NCP | 2020-06-02 | Saint-Luc Bouge* | 104 | 46.7 | 74.3 | 95.8 | 79 | 100 | 100 | 99.81 | √ | Overall sensitivity: | N | 99.632 | TP | % | Overall specificity: | N | 79 | TN | %
[5 empty] | 2020-07-31 | KULeuven* | 223 | 32.6 | 69.4 | 92.9 | 113 | 100 |  |  |  |  | 497 | 207.167 | 456.739 | 91.8992 |  | 248 | 113 | 248 | 100
[5 empty] | 2020-08-18 | AZ Delta* | 170 | 50 | 92.9 | 88.2 | 56 | 100 | [5 empty] | 149.94 |  |  |  |  | 56
15 | Abbott | SARS-CoV-2 IgG assay | CMIA | NCP | 2020-05-13 | UZ Gent* | 203 | 21.2 | 76.6 | 88.4 | 30 | 100 | 100 | 99.63 |  | Overall sensitivity: | N | 179.452 | TP | % | Overall specificity: | N | 30 | TN | %
[5 empty] | 2020-05-11 | Cerba* | 77 | 33.3 | - | 100 | 50 | 100 |  |  |  |  | 765 | 77 | 718.5629 | 93.9298 |  | 440 | 50 | 435.0008 | 98.8638
[5 empty] | 2020-07-31 | KULeuven* | 223 | 27.9 | 67.3 | 95.2 | 113 | 100 | [5 empty] | 212.296 |  |  |  |  | 113
[5 empty] | 2020-05-21 | CHU UCL Namur | 201 | 34.22 | 60.87 | 94.83 | 100 | 97 | [5 empty] | 190.6083 |  |  |  |  | 97
[5 empty] | 2020-06-15 | CHU Liège | 61 | 16.18 | 82.35 | 97.06 | 147 | 98.64 | [5 empty] | 59.2066 |  |  |  |  | 145.0008
16 | Snibe | MAGLUMI 2019-NCoV IgM/IgG Kit | CLIA | ND | 2020-05-10 | UCL, ULB* | 78 | 23 | 48 | 100 | 100 | 100 | 95.6 | 97.33 | √ | Overall sensitivity: | N | 78 | TP | % | Overall specificity: | N | 100 | TN | %
[5 empty] | 2020-07-31 | KULeuven* | 223 | 25.6 | 64.3 | 92.9 | 113 | 96.5 |  |  |  |  | 427 | 207.167 | 395.417 | 92.6035 |  | 285 | 109.045 | 281.045 | 98.6123
[5 empty] | 2020-04-30 | UCL, ULB | 126 | 17.8 | 54.8 | 87.5 | 72 | 100 | [5 empty] | 110.25 |  |  |  |  | 72
17 | Ortho Clinical Diagnostics | VITROS Anti-SARS-CoV-2 Total | CLIA | S1 | 2020-05-21 | CHU UCL Namur | 201 | 42.58 | 69.75 | 96.55 | 100 | 98 | 100 | 100 | √
18 | Bio-Rad | Platelia SARS-CoV-2 Total Ab Assay | ELISA | NCP | 2020-06-28 | HIS-IZZ | 208 | 45 | 87 | 94.7 | 79 | 97.5 | 100 | 100 | √
19 | Diasorin | LIAISON SARS-CoV-2 S1/S2 IgG test | CLIA | S1/S2 | 2020-04-26 | HIS-IZZ | 44 | 13 | 87 | 100 | 81 | 100 | 97.56 | 99.3 | √ | Overall sensitivity: | N | 44 | TP | % | Overall specificity: | N | 81 | TN | %
[5 empty] | 2020-05-11 | Cerba* | 77 | 33.3 | - | 87.5 | 30 | 100 |  |  |  |  | 1125 | 67.375 | 1022.2917 | 90.8704 |  | 602 | 30 | 594.0068 | 98.6722
[5 empty] | 2020-05-15 | AZ Delta* | 168 | 21.95 | 68.89 | 74.7 | 56 | 96.43 | [5 empty] | 125.496 |  |  |  |  | 54.0008
[5 empty] | 2020-06-15 | CHU Liège* | 365 | 31.3 | 70.8 | 100 | 85 | 100 | [5 empty] | 365 |  |  |  |  | 85
[5 empty] | 2020-06-17 | Saint-Luc Bouge* | 137 | 13 | 33.3 | 88.51 | 141 | 97.9 | [5 empty] | 121.2587 |  |  |  |  | 138.039
[5 empty] | 2020-07-07 | LHUB-ULB | 111 | 51.4 | 67.7 | 87.7 | 96 | 97.9 | [5 empty] | 97.347 |  |  |  |  | 93.984
[5 empty] | 2020-07-31 | KULeuven* | 223 | 14 | 58.2 | 90.5 | 113 | 99.1 | [5 empty] | 201.815 |  |  |  |  | 111.983
20 | Eurofins |  | ELISA |  |  | HIS-IZZ
21 | Beijing Wantai Biological Pharmacy Enterprise | WANTAI SARS-CoV-2 Ab ELISA | ELISA | RBD | 2020-05-15 | AZ Delta* | 169 | 61.5 | 95.7 | 92.9 | 55 | 100 | 94.5 | 100 | √
22 | Coris BioConcept | COVID-19 Ag Respi-Strip | rapid | NP |  | ULB, KULeuven, CHU Liège | 156 | 48.3 |  | 100 | 225 | 99.5 | 60.3 | 100 | √
23 | Siemens | SARS-CoV-2 Total (COV2T) (Atellica Solution) | CLIA | RBD | 2020-06-19 | UZA | 170 | 35.7 | 77.8 | 100 | 120 | 100 | 99 | 99 |  | Overall sensitivity: | N | 170 | TP | % | Overall specificity: | N | 120 | TN | %
|  | SARS-CoV-2 Total (COV2T) (Dimension Vista) |  |  |  | GHDC | 39 | - | - | 90 | 40 | 100 |  |  |  |  | 455 | 35.1 | 445.5158 | 97.9156 |  | 242 | 40 | 242 | 100
|  | SARS-CoV-2 Total (COV2T) (Advia Centaur) |  |  | 2020-07-01 | CHU Tivoli | 246 | 46.55 | 85.96 | 97.73 | 82 | 100 | [5 empty] | 240.4158 |  |  |  |  | 82
24 | Xiamen Boson Biotech | RAPID 2019-NCOV IGG/IGM COMBO TEST CARD | rapid | NCP | refused initially | AZ Alma* | 100 | 24.1 | 91 | 100 | 102 | 98 | 87.8 | 99 | √
25 | Vircell | COVID‐19 ELISA IgG | ELISA | S, NCP | 2020-06-23 | AZ Sint-Jan Brugge | 83 | 37.5 | 82.9 | 100 | 185 | 98.92 | 83 | 98 | √
26 | Mikrogen Diagnostik | recomWell SARS-CoV-2 IgG | ELISA | NCP | 2020-05-19 | KULeuven* | 223 | 30.2 | 67.3 | 90.5 | 113 | 96.46 | 100 | 98.7 | √
27 | Surescreen | COVID-19 IgG/IgM Rapid Test Cassette | rapid | ND | 2020-05-19 | KULeuven | 153 | 37.8 | 73.1 | 94.7 | 103 | 95.1 | 91 | 99 | √
28 | Euroimmun | Anti-SARS-CoV-2-NCP ELISA IgG | ELISA | NCP | 2020-05-15 | AZ Delta* | 170 | 55.8 | 90.5 | 80.3 | 56 | 98.2 |  |  | √ | Overall sensitivity: | N | 136.51 | TP | % | Overall specificity: | N | 54.992 | TN | %
[5 empty] | 2020-07-31 | KULeuven | 223 | 30.2 | 71.4 | 95.2 | 113 | 94.7 |  |  |  |  | 530 | 212.296 | 473.202 | 89.2834 |  | 310 | 107.011 | 298.068 | 96.151
[5 empty] | 2020-06-17 | Saint-Luc Bouge* | 137 | 17.4 | 70.4 | 90.8 | 141 | 96.5 | [5 empty] | 124.396 |  |  |  |  | 136.065
29 | Innovita | 2019-nCoV Ab Test (Colloidal Gold) | rapid |  | 2020-05-15 | AZ Delta* | 170 | 35 | 88.9 | 78.8 | 54 | 100 | NO IFU |  | √
30 | AAZ-LMB | COVID-19 IgG/IgM Rapid Test (Whole Blood/Serum/Plasma) – Cassette | rapid | ND | 2020-10-19 | OLVZ Aalst | 81 |  |  | 97 | 85 | 96.5 | 100 | 100 | √
31 | Diapro | COVID-19 IgG | ELISA | NCP, S1 & S2 |  | AZ Alma | [6 empty] | 98 | 98 | √
32 | NG Biotech | NG-Test IgG-IgM COVID-19 All-in-one | rapid | ND |  | ? | [6 empty] | 100 | 100 | √
33 | Abbott | Panbio IgG/IgM Covid-19 rapid test | rapid | ND | 2020-08-03 | Saint-Luc Bouge* | 138 | 34.29 | 80 | 100 | 76 | 98.7 | 95.8 | 94 | √
34 | bioMérieux | VIDAS SARS-CoV-2 IgG | ELFA | ND | 2020-09-11 | AZ Sint-Jan Brugge | 118 | 26.79 | 81.25 | 100 | 180 | 99.44 | 96.6 | 99.9 | √
35 | Intec Products | Rapid SARS-CoV-2 Antibody (IgM/IgG) Test | rapid | ND |  | ? | [6 empty] | 94.4 | 98 | √
36 | Abbott | SARS-CoV-2 IgM Assay | CMIA |  | 2020-10-05 | AZ Sint-Jan Brugge | 143 | 30 | 47.5 | 97.37 | 187 | 98.93 | NO IFU
37 | Euroimmun - See below verified approved after change in sample buffer | Anti-SARS-CoV-2 ELISA (IgA) | ELISA | RBD | 2020-04-26 | HIS-IZZ | 60 | 27 | 100 | 100 | - | ND | 98.6 | 92 | √ | Overall sensitivity: | N | 60 | TP | % | Overall specificity: | N |  | TN | %
[5 empty] | 2020-04-23 | UCL & ULB | 128 | 65.5 | 87.09 | 93.93 | 72 | 86.1 |  |  |  |  | 357 | 120.2304 | 321.8524 | 90.1547 |  | 231 | 61.992 | 189.022 | 81.8277
[5 empty] | 2020-05-19 | KULeuven | - | ND | ND | ND | 103 | 73.8 | [10 empty] | 76.014
[5 empty] | 2020-08-18 | AZ Delta | 169 | 60.4 | 95.2 | 83.8 | 56 | 91.1 | [5 empty] | 141.622 |  |  |  |  | 51.016
38 | Beckman Coulter | Access SARS-CoV-2 IgG | CLIA | RBD | 2020-09-02 | Noorderhart Pelt | 130 | 34 | 77.78 | 100 | 49 | 100 | 100 | 99.8 | √
39 | Meso Scale Discovery |  |  |  |  | Gent or UA ?
40 | Theradiag | ELISA COVID-19 THERA02 IgG | ELISA | S |  | OLVZ Aalst | 81 |  |  | 87.9 | 85 | 98.8 | 100 | 100 | √
41 | Theradiag | TDR Covid-19 IgG+IgM THERA | rapid |  |  | OLVZ Aalst | 81 |  |  | 90.9 | 85 | 82.4 | 94.7 | 98.7 | v
42 | ID.Vet | ID Screen SARS-CoV-2-N IgG Indirect | ELISA | NCP |  | OLVZ Aalst | 81 | 36.7 | 84 | 93.9 | 85 | 98.8
43 | NovaTec Immundiagnostica | NovaLisa SARS-CoV-2 (COVID-19) IgG | ELISA | NCP | 2020-05-31 | CHU UCL Namur | 201 | 42.1 | 58.1 | 87.9 | 100 | 97 | 100 | 99.2 |  | Overall sensitivity: | N | 176.679 | total | % | Overall specificity: | N | 97 | TN | %
[5 empty] | 2020-06-28 | HIS-IZZ | 208 | 35 | 75 | 94.9 | 79 | 96.2 |  |  |  |  | 409 | 197.392 | 374.071 | 91.4599 |  | 179 | 75.998 | 172.998 | 96.6469
44 | NovaTec Immundiagnostica | NovaLisa SARS-CoV-2 (COVID-19) IgA | ELISA | NCP | 2020-06-28 | HIS-IZZ | 208 | 35 | 85 | 89.7 | 79 | 98.7 | ND | ND
45 | NovaTec Immundiagnostica | NovaLisa SARS-CoV-2 (COVID-19) IgM | ELISA | NCP | 2020-06-28 | HIS-IZZ | 208 | 15 | 55 | 48.7 | 79 | 98.7 | 57.1 | 100
46 | Vircell | COVID-19 ELISA IgM+IgA | ELISA | S, NCP | 2020-06-26 | AZ Sint-Jan Brugge | 81 | 68.75 | 92.86 | 100 | 184 | 89.1 | 88 | 99 | √
47 | Shenzhen YHLO Biotech | iFlash-SARS-CoV-2 IgM | CLIA | ND | 2020-04-29 | Eupen | 108 | 22 | 60 | 63.6 | 75 | 98.7 | 86.1 | 99.2 | √
48 | PRIMA Lab | COVID-19 IgG/IgM RAPID TEST (new) | rapid | NCP |  |  | 99 |  |  |  | 180 | 98.3 |  |  | √ | Sensitivity was calculated for each antibody using 2 different reference methods => PCR: IgG 95% IgM 65.7%. ELISA: IgG 98.92% IgM 92.98%. No validation of specimen.
49 | Ortho Clinical Diagnostics | VITROS Anti-SARS-COV-2 IgG | CLIA | S | 2020-09-10 | CHU UCL Namur | 67 | 33.3 | 70 | 100 | 40 | 97.5 | 90 | 100
50 | bioMérieux | VIDAS SARS-COV-2 IgM | ELFA |  |  | AZ Sint-Jan Brugge | 173 |  |  | 94.59 | 185 | 98.38
51 | Diasorin | IgM
52 | Multi-G | MGS (new)
53 | Prestige Diagnostics | 2019-nCoV IgG/IgM WB Device | rapid | ND |  |  | 226 |  |  | 93.8 | 470 | 96.8 |  |  | √ | No interference study was made, no declaration of conformity
54 | Siemens | SARS-CoV-2 IgG Assay (Dimension Vista) | CLIA |  |  | GHDC | 39 |  |  | 87 | 40 | 100 | 100 | 100
55 | ATLAS Medical | COVID-19 IgG/IgM Rapid Test Cassette (Whole Blood /Serum/Plasma) | rapid | ND |  |  | 264 | 40 | 87.5 | 97.3 | 213 | 98.6
56 | Mylan
57 | Phadia | Sars-CoV-2 Spike 1 IgG antibody test | ELFA | S1 |  |  | 24 |  |  | 100 | 747 | 100 |  |  | √ | No validation of specimen. 771 samples but which specimen?Only 24 samples positive samples. No interference study. No labels.
58 | Roche | Elecsys Anti-SARS-CoV-2 S | ECLIA | S |  |  | 1423 |  |  | 98.8 | 5991 | 99.98 |  |  | √ | No validation of specimen. 7414 samples but which specimen?
59 | Biomerica | COVID-19 IgG/IgM Rapid Strip Test | rapid | NCP |  |  | 64 |  |  | 96.9 | 294 | 96.6 |  |  | √ | 358 samples but which specimen? Waiting for updated ISO certificate from NB. No validation of specimen.
60 | Nantong Egens Biotechnology | Covid-19 IgG/IgM Rapid Test | rapid | ND | [11 empty] | No performance data available
61 | Syntron Bioresearch | QuikPac II COVD-19 IgG/IgM Rapid Test | rapid | NCP |  |  | 76 |  |  | 89.5 | 526 | 97.7 |  |  | √ | 602 samples but which specimen?
62 | The Binding Site Group | Human anti-IgG/A/M SARS-CoV-2 ELISA | ELISA | S |  |  | 162 |  |  | 96.3 | 707 | 99.3 |  |  | √ | Labelling not clear enough
63 | D-tek | BlueDiver COVIDOT 5 IgG | ELISA | NCP, S, S1, S2, RBD |  |  | 218 |  |  | 93.6 | 118 | 90.7 |  |  | √ | No data available for plasma, no validation of specimen
64 | Diesse Diagnostica | CHORUS SARS-CoV-2 IgG | ELISA | ND |  |  | 207 |  |  | 100 | 330 | 97.9 |  |  | √ | Specificity in IFU = 99.6% but recalculated with cross reactivity data included
65 | Theradiag | COVID19-LISA IgG | ELISA | NCP, S |  | OLVZ Aalst | 81 |  |  | 90.9 | 85 | 96.5
66 | SD Biosensor (distributed by Roche) | SARS-CoV-2 Rapid Antibody Test | rapid | NCP, S |  |  | 103 |  | 92.57 | 99.03 | 222 | 98.65
67 | Abbott | SARS-CoV-2 IgG II Quant Architect | CLIA |  |  |  | 158 |  |  | 99.37 | 2008 | 99.55 |  |  | √
68 | Abbott | SARS-CoV-2 IgG II Quant Alinity | CLIA |  |  |  | 155 |  |  | 99.35 | 2008 | 99.6 |  |  | √
69 | Euroimmun | anti-SARS-CoV-2 QuantiVac IgG ELISA | ELISA | S1 |  |  | 105 |  |  | 96.3 | 1458 | 99.8 |  |  | √
70 | Alphadia | BlueDiver COVIDOT 5 IgG | ELISA | NCP, S, S1, S2, RBD |  |  | 218 |  |  | 93.6 | 118 | 91 |  |  | √
71 | ZENUM Sarl | COVID-19 IgG/IgM Rapid Test Cassette (WB/S/P) | rapid |  |  |  | 37 | [7 empty] | √ | Sensitivity in IFU: 97% (IgM - 37 stalen) en 97.4% (IgG - 77 stalen) en Specificiteit: 100% (IgM - 40 stalen)
72 | Diasorin | LIAISON SARS-CoV-2 TrimericS IgG | CLIA | S |  |  | 155 |  |  | 98.7 | 1899 | 99.5 |  |  | √
73 | Hangzhou York Biotech | SARS-CoV-2 IgM/IgG Rapid Test Kit | rapid |  |  |  | 163 |  |  | 99.4 | 153 | 100 |  |  | √
74 | Siemens Healthineers | Atellica IM SARS-CoV-2 IgG (sCOVG) | CLIA |  |  |  | 79 |  |  | 91.14 | 1995 | 99.9 |  |  | √
75 | Euroimmun | anti-SARS-CoV-2 ELISA (IgA) | ELISA | S1 |  |  | 137 |  | 88.2 | 96.9 | 210 | 98.3 | 96.9 | 98.3 | √
76 | Diesse Diagnostica Senese | SARS CoV-2 SCREEN Serum | ELISA | total Ab |  |  | 86 |  |  |  | 292 |  | 98.8 | 100 | √
77 | ACON Biotech (Hangzhou) | SARS-CoV-2 IgG/IgM Rapid Test | rapid | ND |  |  | 359 |  |  |  | 820 |  | 98.89 | 98.05 | √
78 | BioMARIC | SARS-CoV-2 Total Ab ELISA | ELISA | RBD |  |  | 126 |  |  | 99.21 | 352 | 100 |  |  | √
[7 empty] | 14991 | AVERAGE | 74.3202 | 93.4191 | 27837 | 97.157 | 94.65 | 98.763
[8 empty] | MEDIAN | 73.8 | 94.7 |  | 98.6 | 97.56 | 99.2
[7 empty] | MIN | 48.7
[7 empty] | 1st quart | 89.65 | 40.95
[7 empty] | mediaan | 94.315 | 4.665
[7 empty] | 3rd quart | 97.4 | 3.085
[7 empty] | max | 100 | 2.6

## Website Publication Table
List of recommended SARS-CoV-2 antibody and antigen tests
Latest update: 06/04/2022
Validated antibody tests
Manufacturer | Test | Type | Target | Validation link | Qualitative Semi-quantitative Quantitative
AAZ-LMB | COVID-19 IgG/IgM Rapid Test (Whole Blood/Serum/Plasma) – Cassette | Rapid |  | to be published | qualitative
Abbott | Panbio IgG/IgM Covid-19 rapid test | Rapid |  | to be published | qualitative
Avioq | Novel Coronavirus (2019-nCoV) Antibody IgG/IgM Assay (Colloidal Gold) | Rapid |  | Publication | qualitative
BMT Bio marketing | Lab on Time 2019-nCoV IgG/IgM rapid test cassette | Rapid |  | Publication | qualitative
Dynamiker Biotechnology (Tianjin) | 2019 nCOV IgG/IgM Rapid Test | Rapid | NCP | Publication | qualitative
Liming Bio | StrongStep SARS-CoV-2 lgM/lgG Antibody Rapid Test | Rapid |  | Publication | qualitative
Surescreen | COVID-19 IgG/IgM Rapid Test Cassette | Rapid |  | Publication | qualitative
VivaChek | VivaDiag SARS-CoV-2 IgM/IgG Rapid Test | Rapid |  | Publication | qualitative
Xiamen Boson Biotech | RAPID 2019-NCOV IGG/IGM COMBO TEST CARD | Rapid | NCP | Report | qualitative
Zentech | quickZen COVID-19 IgM/IgG rapid test | Rapid | RBD | Publication | qualitative
Abbott | SARS-CoV-2 IgG Assay (Alinity, Architect) | CMIA | NCP | to be published | semi-quantitative
| SARS-CoV-2 IgM Assay (Alinity, Architect) | CMIA |  | Report | semi-quantitative
Beckman-Coulter | Access SARS-CoV-2 IgG | CLIA | RBD | Report | semi-quantitative
Beijing Wantai Biological Pharmacy Enterprise | WANTAI SARS-CoV-2 Ab ELISA | ELISA | RBD | to be published
bioMérieux | VIDAS SARS-CoV-2 IgG | ELFA |  | Report | semi-quantitative
| VIDAS SARS-CoV-2 IgM | ELFA |  | Report | semi-quantitative
Bio-Rad | Platelia SARS-CoV-2 Total Ab Assay | ELISA | NCP | Publication | semi-quantitative
Diasorin | LIAISON SARS-CoV-2 S1/S2 IgG test | CLIA | S1/S2 | Publication | quantitative
Euroimmun | Anti-SARS-CoV-2 ELISA IgG | ELISA | S1 | Publication 1 | semi-quantitative
Publication 2
Euroimmun | Anti-SARS-CoV-2-NCP ELISA IgG | ELISA | NCP |  | semi-quantitative
ID.Vet | ID Screen SARS-CoV-2-N IgG Indirect | ELISA | NCP | to be published | semi-quantitative
Mikrogen Diagnostik | recomWell SARS-CoV-2 IgG | ELISA | NCP | to be published | quantitative
NovaTec Immundiagnostica | NovaLisa SARS-CoV-2 (COVID-19) IgG | ELISA | NCP | Publication | semi-quantitative
Ortho Clinical Diagnostics | VITROS Anti-SARS-COV-2 IgG | CLIA | S | to be published | semi-quantitative
| VITROS Anti-SARS-CoV-2 Total | CLIA | S1 | to be published | semi-quantitative
Roche | Elecsys Anti-SARS-CoV-2 | ECLIA | NCP | Publication | qualitative
Shenzhen YHLO Biotech | iFlash-SARS-CoV-2 IgG | CLIA |  | Report
Siemens | SARS-CoV-2 Total (COV2T) (Advia Centaur, Atellica Solution, Dimension Vista) | CLIA | RBD | to be published
Snibe | MAGLUMI SARS-CoV-2 IgM/IgG Test | CLIA |  | Publication | quantitative
Theradiag | COVID19-LISA IgG | ELISA | NCP, S | to be published | semi-quantitative
Vircell | COVID‐19 ELISA IgG | ELISA | S, NCP | Report | semi-quantitative
Verified antibody tests, according to the new procedure (not validated by Belgian lab).
Please note: The new procedure can be found at: https://www.fagg.be/sites/default/files/content/new_application_procedure.pdf
ACON Biotech (Hangzhou) | SARS-CoV-2 IgG/IgM Rapid Test | Rapid |  | NA | qualitative
ATLAS Medical | COVID-19 IgG/IgM Rapid Test Cassette | Rapid |  | NA | qualitative
Beckman-Coulter | Access SARS-CoV-2 IgG (1st IS) | CLIA | RBD | NA | qualitative + quantitative
Biomerica | COVID-19 IgG/IgM Rapid Strip Test | Rapid | NCP | NA | qualitative
Hangzhou York Biotech | SARS-CoV-2 IgM/IgG Rapid Test Kit | Rapid |  | NA | qualitative
Prestige Diagnostics | 2019-nCoV IgG/IgM WB Device | Rapid |  | NA | qualitative
PRIMA Lab | COVID-19 IgG/IgM RAPID TEST (new) | Rapid | NCP | NA | qualitative
SD Biosensor (distributed by Roche) | SARS-CoV-2 Rapid Antibody Test | Rapid | NCP, S | NA | qualitative
ZENUM | COVID-19 IgG/IgM Rapid Test Cassette (WB/S/P) | Rapid |  | NA | qualitative
Abbott | SARS-CoV-2 IgG II Quant (Alinity, Architect) | CMIA | RBD | NA | quantitative
Alphadia | Covid-19 Profile 5 IgG DOT | ELISA | NCP, S, S1, S2, RBD | NA | qualitative
BioMARIC | SARS-CoV-2 Total Ab ELISA | ELISA | RBD | NA | qualitative
Diasorin | LIAISON SARS-CoV-2 TrimericS IgG | CLIA | S | NA | quantitative
Diesse Diagnostica Senese | CHORUS SARS CoV-2 Screen Serum | ELISA |  | NA | qualitative
D-tek | BlueDiver COVIDOT 5 IgG | ELISA | NCP, S, S1, S2, RBD | NA | qualitative
Euroimmun | Anti-SARS-CoV-2 QuantiVac ELISA (IgG) | ELISA | S1 | NA | quantitative
| Anti-SARS-CoV-2-ELISA (IgA) | ELISA | S1 | NA | semi-quantitative
Phadia | EliA SARS-CoV-2-Sp1 IgG | ELFA | S1 | NA | quantitative
Roche | Elecsys Anti-SARS-CoV-2 S | ECLIA | S | NA | quantitative
Siemens | SARS-CoV-2 IgG Assay | CLIA | S1 | NA | semi-quantitative
| Atellica IM SARS-CoV-2 IgG (sCOVG) | CLIA |  | NA | quantitative
The Binding Site Group | Human anti-IgG/A/M SARS-CoV-2 ELISA | ELISA | S | NA | qualitative
Antigen tests for professional use with sensitivity ≥ 90% AND specificity ≥ 99% as claimed in the IFU.
Please note: At the time this list was compiled, extensive independent validations were scarce for most SARS-CoV-2 antigen tests. Therefore, the sensitivity and specificity values in the list below are taken from the manufacturer’s instructions for use and should be interpreted with caution. Antigen test show a large range in sensitivity values that are lower than RT-PCR sensitivity. Real-life performance of antigen tests depends on several factors such as sample type, specimen collection and handling, infection stage and certainly viral load. Antigen tests perform best when the person is tested in the early stages of infection with SARS-CoV-2 when viral load is generally highest. More information on antigen test validation studies can for instance be found at The Foundation for Innovative New Diagnostics (FIND) that is conducting prospective diagnostic evaluation studies in collaboration with multiple, independent sites to determine the accuracy of COVID-19 antigen rapid tests. Precise knowledge of testing circumstances, test performance characteristics and specific patient’s clinical conditions are imperative for a correct interpretation of antigen test results. The tests on this list are for professional use only and cannot be used or sold as self-tests. If a test is also taken up in the common Health Security Committee (HSC) list (colum G), it can be used in the context of the EU Digital Green Certificate ('test certificate').
Manufacturer | Test | Sensitivity | Specificity | Specimen type | Test ID in JRC database | Test on common HSC list?
AAZ | COVID-VIRO | 96.6 | 100 | NP swab | 1833 | Yes
Abbott | PANBIO COVID-19 Ag rapid test device | 98,1/93,3 | 99,8/99,4 | Nasal swab/NP swab | 1232 | Yes
ACON Biotech (Hangzhou) | Flowflex SARS-CoV-2 Antigen Rapid Test | 97.1 | 99.4 | Nasal swab/NP swab | 1468 | Yes
ACON Laboratories |  |  |  |  | 1468 | Yes
ACRO Biotech | COVID-19 Antigen Rapid Test (Oral Fluid) | 90.3 | 99.3 | Saliva | 1949 | No
| SARS-CoV-2 Antigen Rapid Test (Swab) | 94.6 | 99.4 | Nasal swab | 1688 | No
AMEDA Labordiagnostik | AMP Rapid Test SARS-CoV-2 Ag | 97.3 | 100 | NP swab | 1304 | Yes
AMEDA Labordiagnostik | AMP Rapid Test SARS-CoV-2 Ag - INFLUENZA A+B | 97.3 | 100 | NP swab | -
Anhui Deepblue Medical Technology | COVID-19 (SARS-CoV-2) Antigen Test Kit (Colloidal Gold) | 95.7/96.4 | 99.3/99.8 | NP swab/OP swab | 1736 | Yes
| COVID-19 (SARS-CoV-2) Antigen Test Kit (Colloidal Gold) | 96.4/97.1 | 99.8/99.8 | Nasal swab/Saliva | 1589 | No
Anhui Formaster Biosci Co | New Coronavirus (COVID-19) Antigen Rapid Test (Saliva) | 96.24 | 99.8 | Saliva | 2480 | No
Ardi Research | CoVard COVID-19 Antigen Rapid Test | 96.66 | 98.96 | NP swab | - | No
Asan Pharmaceutical Co | Asan Easy Test Flu/COVID-19 Ag Combo | 94.7 | 99.1 | NP swab | 1761 | No
Assure Tech (Hangzhou) Co | COVID-19 Antigen Rapid Test Device | 94.3/91.6 | 99.0/100 | NP swab/Nasal swab | 2270 | No
| COVID-19 Antigen Nasal Test Kit | 94.5 | 100 | Nasal swab | 2076 | No
Asterion Otel İnşaat Bilişim Medikal Maden Tic.Ltd.Şti. | As-check | 95.7 | 99.2 | NP swab | 2454 | No
ATLAS Medical | Covid-19 Antigen Rapid Test Cassette | 92.4 | 99.1 | NP swab | - | No
BBB Inc. | MARK-B™ COVID-19 Ag | 91.67 | 99.59 | NP swab | 2315 | No
Beijing Beier Bioengineering | COVID-19 Antigen Rapid Test Kit | 96.5 | 99.7 | Nasal swab/NP swab/OP swab | 1310 | No
Beijing Hotgen Biotech | Novel coronavirus 2019 n-CoV Ag Test | 98,6/97,3 | 100/99,2 | Nasal swab/OP swab | 1870 | Yes
Beijing Lepu Medical | SARS-CoV-2 Antigen Rapid Test Kit | 92 | 99.3 | Nasal swab | 1331 | Yes
Bioscience (Tianjin) Diagnostic Technology | Novel Coronavirus (2019-nCoV) Antigen Rapid Detection kit | 95.83 | 99.31 | Nasal swab | 2623 | No
Biospeedia | Covid-19 Nasal Ag Test | 95 | 99.6 | Nasal swab | 2503 | No
Biospeedia | COVID-19 Speed-Antigen Test | 97.46 | 99.28 | NP swab | 2380 | No
Biosynex | BIOSYNEX COVID-19 Ag BSS | 96 | 100 | NP swab | 1223 | Yes
Bioteke Corporation (Wuxi) Co | SARS-CoV-2 Antigen Test Kit (Colloidal gold method) | 96.49 | 99.28 | NP swab | 2067 | Yes
Biotical Health | SARS-CoV-2 Ag CARD | 92.9 | 99.6 | NP swab | 2013 | Yes
Canary Global Inc | SARS-CoV-2 Ultra Rapid Mobile Antigen Test | 96.4/97.30 | 99.17/99.17 | saliva/NP swab | 2283 | No
CerTest BIOTEC | One Step SARS-CoV-2 AG card | 92.9 | 99.6 | NP swab | 1173 | Yes
Changzhou Biowin Pharmaceutical Co | Novel Coronavirus (COVID-19) Antigen Test Kit (Colloidal Gold) | 95.6 | 99 | Nasal swab | - | No
Chil Tibbi Mal. San. Tic. | CHIL COVID-19 Antigen Rapid Test | 99.0/99.0/98.0 | 99.6/99.6/99.6 | NP swab/OP swab/Nasal swab | 1691 | Yes
Chongqing M&D Biotechnology Co | 2019-nCoV Antigen Test Kit | 91.53 | 100 | NP swab | 1674 | No
Citest Diagnostics | COVID-19 Antigen Rapid Test (Swab) | 96.4/92.9 | 99.0/100 | NP swab/Nasal swab | 1387 | No
Coris BioConcept | Covid-19 Ag Respi-strip with reference C-1123 (for dry swab procedure only) | 91.2 | 99.4 | NP swab | 133 | No
Coris BioConcept | COVID-19 Ag K-SeT | 96.9 | 100 | Nasal swab | 2648 | No
Cypress Diagnostics | COVID-19 Ag Rapid Test | 93 | 100 | NP swab | - | No
Dialab | Diaquick Covid-19 Ag Cassette Z20401CE | 93.2 | 100 | NP swab | 1375 | No
| Diaquick Covid-19 Ag Cassette Z20601CE | 96.4 | 99.2 | NP swab
DiaSorin | LIAISON SARS-CoV-2 Ag | 97,1/94,6 | 100,0/99,5 | NP swab/Nasal swab | 1960 | Lab test
DiaSource | Covid-19 Antigen Rapid Test | 96.4 | 99.2 | NP swab | 1792 | No
Dräger Safety AG & Co KGaA | Antigen Test SARS-CoV-2 | 92.3 | 99.6 | Nasal swab | 2273 | Yes
Dynamiker | Dynamiker SARS-CoV-2 Ag Rapid Test | 95 | 99.1 | Nasal swab/OP swab | 1250 | No
Eon Biotechnology | EONbt COVID-19 Antigen | 94.89 | 99.25 | Nasal swab | - | No
Generic Assays | GA CoV-2 Antigen Rapid 3980 | 96.4 | 99.2 | Nasal swab | 1580 | No
Genobio Pharmaceutical Co | SARS-CoV-2 Antigen Rapid Test (Colloidal Gold) | 97.3/96.30/97.14/95.24 | 99.14/99.28/99.34/99.26 | NP swab/nasal swab/OP swab/saliva | 2642 | Yes
Genrui Biotech | SARS-CoV-2 Antigen Test Kit (Colloidal Gold) | 95.03 | 99.02 | NP swab/OP swab | 2012 | Yes
Getein Biotech | One Step Test for SARS-CoV-2 Antigen (Colloidal Gold) | 97.1 | 99 | Nasal swab | 2183 | Yes
Goldsite Diagnostics | SARS-CoV-2 Antigen Kit (Colloidal Gold) | 93.04 | 100 | Nasal swab | 1197 | Yes
Green Cross Medical Science Corp. | Genedia W Covid-19 Ag | 90.2 | 100 | NP swab | 1144 | Yes
Guangdong Hecin Scientific | 2019-nCoV Antigen Test Kit (Colloidal Gold method) | 96.83 | 99.39 | NP swab | 1747 | Yes
Guangdong Longsee Biomedical | 2019-nCoV Ag Rapid Detection Kit (Immuno-Chromatography) | 95.2 | 99.7 | NP swab | 2504 | No
| 2019-nCoV Ag Rapid Detection Kit (Immuno-Chromatography) | 94.7 | 99.7 | OP swab | 2509 | No
| 2019-nCoV Ag Rapid Detection Kit (Immuno-Chromatography) | 92.4 | 99.7 | Saliva | 1216 | Yes
Guangzhou Wondfo Biotech | Wondfo SARS-CoV-2 Ag Test | 96.18 | 99.72 | NP swab/OP swab | 1214 | No
Hangzhou Alltest Biotech | COVID-19 Antigen Rapid Test (Nasopharyngeal Swab) | 90.4 | 99.5 | NP swab | 1257 | Yes
| COVID-19 Antigen Rapid Test (Oral Fluid) | 90.1 | 99.3 | Saliva | 1947 | No
| SARS-CoV-2 Antigen Rapid Test (Nasal Swab) | 93.4 | 100 | Nasal swab | 1946 | No
| SARS-CoV-2 Antigen Rapid Test | 93.4/93.5 | 99.9/99.3 | NP swab/nasal swab | 2384 | No
Hangzhou Biotest Biotech | SARS-CoV-2 Antigen Rapid Test Cassette | 95 | 99.2 | NP swab | 1868 | No
Hangzhou Clongene Biotech | Lungene Covid-19 Ag Rapid Test | 91.4 | 100 | NP swab/OP swab | 1610 | Yes
Hangzhou Fanttest Biotech | Novel Coronavirus (SARS-CoV-2) Antigen Rapid Test Cassette | 98.08/98.06 | 99.34/99.45 | NP swab/nasal swab | 2153 | No
Hangzhou Genesis Biodetection and Biocontrol | KaiBiLi COVID-19 Antigen | 96.6 | 100 | Nasal swab | 2085 | No
| KaiBiLi COVID-19 Antigen Neo | 96.6 | 100 | Nasal swab | - | No
Hangzhou Laihe Biotech Co | LYHER Novel Coronavirus (COVID-19) Antigen Test Kit (Colloidal Gold) | 100 | 99.7 | OP swab | 2224 | No
Hangzhou Realy Tech | Novel coronavirus (SARS-CoV-2) Antigen Rapid Test Cassette (swab) | 96.17 | 100 | NP swab/OP swab | 1220 | No
| Novel Coronavirus (SARS-Cov-2) Antigen Rapid Test Device | 92.9 (Ct<30) | 99.58 | Saliva | 1796 | No
| Novel Coronavirus (SARS-Cov-2) Antigen Rapid Test Device (nasal swab) | 95.4 | 100 | Nasal swab | 2018 | No
Hangzhou Sejoy Electronics & Instruments | SARS-CoV-2 Antigen Rapid Test Cassette | 97.4/97.4/95.8/95.7 | 99.1/99.1/99.4/99.3 | Nasal swab/NP swab/OP swab/Saliva | 1952 | No
Hangzhou York Biotech | SARS-CoV-2 Antigen Rapid Test Kit | 94.59 | 100 | NP swab | - | No
Healgen Scientific (distributed by Siemens) | CLINITEST Rapid COVID-19 Antigen Test | 98,3/97,3 | 99,6/100 | NP swab/Nasal swab | 1218 | Yes
Hubei Jinjian Biology Co | SARS-CoV-2 Antigen Test Kit (Colloidal Gold Method) | 97.5/97.5 | 99.2/99.6 | NP swab/OP swab | 1759 | Yes
Humasis | Covid-19 Ag Rapid Test | 95.5 | 100 | NP swab | 1263 | Yes
InTec products Inc | Rapid SARS-CoV-2 Antigen Test | 96.3/95.5 | 100/99.6 | NP swab/Nasal swab | 1783 | No
InTec products Inc | Rapid SARS-CoV-2 Antigen Test | 93.1 | 100 | NP swab | 2419 | Yes
iXensor Co | PixoTest COVID-19 Ag Test Kit | 94.17 | 99.37 | Nasal swab | 2217 | No
Jiangsu Diagnostics Biotechnology Co | COVID-19 Antigen Rapid Test Cassette (Colloidal gold) | 97.58/97.92 | 100 | NP swab/OP swab | 1920 | Yes
Jiangsu Konsung Bio-Medical Science and Technology | COVID-19 Antigen Rapid Test Kit (Colloidal Gold) | 97.5/98.3/96.0 | 99.1/99.2/99.5 | Nasal swab/NP swab/OP swab | 1899 | No
| COVID-19 Salivary Antigen Rapid Test Kit (Colloidal Gold) | 96.3 | 99.34 | Saliva | 1924 | No
Jiangsu Medomics Medical Technology | SARS-CoV-2 Antigen Test Kit (LFIA) | 97.73/95.54 | 99.51/99.04 | NP swab+ OP swab/nasal swab | 2006 | Yes
Jinan Babio Biotechnology Co | SARS-CoV-2 Antigen Rapid Detection Kit (Colloidal Gold) | 96.38/97.41/97.48 | 100/99.53/99.65 | Nasal swab/OP swab/saliva | 2151 | No
JOYSBIO (Tianjin) Biotechnology | SARS-CoV-2 Antigen Rapid Test Kit (Colloidal Gold) | 95.1 | 100 | Saliva | 1764 | Yes
LiClear Biotech (Hangzhou) | SARS-COV-2 Nucleocapsid (N) Antigen Rapid Test Cassette | 95.1 | 99.3 | NP swab | - | No
LumiQuick Diagnostics | QuickProfile COVID-19 Antigen Test Strip | 94 | 99 | NP swab | 1267 | Yes
MEDsan | SARS-CoV-2 Antigen Rapid Test | 92.5 | 99.8 | Nasal swab/OP swab | 1180 | Yes
MP Biomedicals | Rapid SARS-CoV-2 Antigen Test Card | 96.4 | 99.03 | NP swab/OP swab | 1481 | Yes
Multi-G | Covid19CHECK-GEN | 92.5 | 99.2 | NP swab/OP swab | - | No
| COVID19CHECK-NAS (COVID-19 antigen rapid test cassette) | 98.1 | 99.2 | Nasal swab | 2260 | Yes
| COVID19CHECK-SAL (Sars-Cov-2 Antigen Rapid Test) | 96.8 (Ct≤ 33) | 99.1 | Saliva | - | No
Mylab Discovery Solutions Pvt | PathoCatch COVID-19 Antigen Lateral Flow Test Device | 92 | 100 | Nasal swab | - | No
nal von minden | NADAL COVID-19 Ag Rapid Test | 97.56 | 99.9 | NP swab/OP swab | 1162 | Yes
Nanjing Synthgene Medical Technology | SARS-COV-2 Nucleocapsid (N) Antigen Rapid Detection Kit (Colloidal gold method) | 99.33 | 99.5 | NP swab | 2164 | No
Nanjing Vazyme Medical Technology | Severe Acute Respiratory Syndrome Coronavirus 2 (SARS-CoV-2) Antigen Detection Kit (Colloidal Gold-Based) | 97.6 | 99.3 | Nasal swab/OP swab | 1849 | No
NanoEntek | FREND COVID-19 Ag | 94.7 | 100 | NP swab | 1420 | Yes
Nantong Diagnosis Biotechnology | Covid-19 Antigen Saliva Test Kit (Colloidal Gold) (Cassette) | 98.1 | 99.3 | Saliva | 1617 | No
| Covid-19 Antigen Test Kit (Colloidal Gold) | 98,8/98,5 | 100 | NP swab/OP swab | - | No
New Gene (Hangzhou) Bioengeneering | Covid-19 Antigen Detection Kit | 97.3/95.7/98/95.1 | 99/99/99.1/99.1 | Nasal swab/OP swab/NP swab/Sputum | 1501 | Yes
ONCOSEM Onkolojik Sistemler San. ve Tic. A.Ş. | 2019-nCoV Antigen Rapid Test Kit | 97.6 | 99.3 | NP swab | - | No
Ortho-Clinical Diagnostics | VITROS SARS-CoV-2 Antigen test | 97.8 | 99.2 | NP swab | 1200 | Lab test
PCL | PCL COVID19 Ag Gold | 90.8 (Ct<30)/91.7 | 99.5/100 | Saliva/NP swab | 1758 | No
Prestige Diagnostics | 2019-nCoV Antigen Device | 90.9 | 99.1 | NP swab | - | No
Prestige Diagnostics | 2019-nCoV Antigen Device (Anterior Nasal Swab) | 96.3 | 99.5 | Nasal swab | 1202 | No
PRIMA Lab | PRIMACOVID COVID-19 Antigen Rapid Test | 96.4/92.9 | 99.2/100 | NP swab/Nasal swab | 1797 | No
Qingdao Hightop Biotech | SARS-CoV-2 Antigen Rapid Test (Immunochromatography) | 92.7/95.0/95.0 | 99.8/99.8/99.8 | Nasal swab/NP swab/OP swab | 1341 | Yes
Quidel | Sofia SARS Antigen FIA | 96.7 | 100 | NP swab/Nasal swab | 1097 | Yes
| Sofia 2 Flu + SARS Antigen FIA | 95.2 | 100 | NP swab/Nasal swab | - | No
Roche Diagnostics | Elecsys SARS-CoV-2 Antigen | 94.5 | 99.9 | NP swab/OP swab | 2156 | Lab test
SD Biosensor (distributed by Roche) | SARS-CoV-2 Rapid Antigen Test | 96.52 | 99.68 | NP swab | 1604 | Yes
Shenzhen Huian Biosci Technology | SARS-CoV-2 Antigen Test Kit (Colloidal Gold) | 94.6 | 99.1 | Nasal swab | - | No
Shenzhen Landwind Biotechnology | SARS-CoV-2 Antigen Test Kit (Colloidal Gold) | 95.7 | 99.2 | Nasal swab | - | No
Shenzhen Lifotronic Technology Co | SARS-CoV-2 Antigen (GICA) | 96.38 | 99.38 | Nasal swab | 2148 | No
Shenzhen Lvshiyuan Biotechnology | Green Spring SARS-CoV-2-Antigen-Rapid Test Kit (Colloidal Gold) | 96.8/98 | 100 | Nasal swab/NP swab | 2109 | Yes
Shenzhen Microprofit Biotech | Fluorecare SARS-CoV-2 Test Kit | 98.6/100 | 100 | NP swab/Nasal swab | 1967 | Yes
| Fluorecare SARS-CoV-2 Spike Protein Test Kit | 98.6 | 100 | NP swab | 1228 | Yes
Shenzhen Superbio Technology Co | 2019-nCoV Antigen Rapid Test (Immunochromatography Assay) | 96.5/93.8/93.8 | 99.7/100/99.4 | Nasal swab/NP swab/OP swab | - | No
Shenzhen Ultra Diagnostics Biotech | SARS-CoV-2 Antigen Test kit (SC0201 + SC0202) | 92/100/96 | 100 | NP swab/OP swab/Saliva | - | No
| SARS-CoV-2 Antigen Test kit (SC0203 + SC0204) | 92/95,7/97,3 | 100/99/99 | NP swab/OP swab/Sputum | - | No
Shenzhen YHLO Biotech Co | GLINE-2019-nCoV Ag | 96.5/97.4 | 99.3/99.3 | NP swab/Nasal swab | 1347 | Yes
Surescreen | COVID-19 Coronavirus Rapid Antigen Test Cassette | 93.3 | 100 | NP swab/OP swab | 1275 | No
Todapharma | TODA Coronadiag Ag | 96.6 | 100 | NP swab/OP swab | 1466 | Yes
Ulti med Products | COVID-19 Antigen Test (Nasopharyngeal Swab) | 96.4 | 99.2 | NP swab | 1455 | No
| COVID-19 and Influenza A+B Antigen Combo Rapid Test (Nasopharyngeal Swab) | 96.4 | 99.2 | NP swab | 1391 | No
Van Oostveen Medical | Coronavirus Ag Rapid Test Cassette (Swab) | 96.7 | 99.2 | NP swab | 1831 | No
Vision Biyoteknoloji | Vision Covid-19 Rapid Antigen Test kit | 97.57 | 99.6 | Nasal swab | 2776 | No
Vitrosens Biyoteknoloji Ltd. Şti. | RapidFor SARS-CoV-2 Rapid Antigen Test Kit | 97.3/95.3/97.05 | 99.05/100/99.23 | NP swab/OP swab/Nasal swab | 1443 | Yes
VivaChek Biotech | VivaDiag Pro SARS-CoV-2 Ag Rapid Test | 96.1/96.1/97.0 | 100 | NP swab/OP swab/Nasal swab | 2103 | No
| VivaDiag SARS-CoV-2 Ag Rapid Test | 95 | 100 | NP swab/OP swab/Nasal swab | 2102 | No
| VivaDiag SARS-CoV-2 Ag Saliva Rapid Test | 98.3 | 100 | Saliva | 1980 | No
Wenzhou OJA Biotechnology | SARS-CoV-2 Antigen Detection Kit (Colloidal Gold Method) | 97.6 | 99.2 | Nasal swab | 1575 | No
Wuhan UNscience Biotechnology Co | SARS-CoV-2 Antigen Rapid Test Kit | 96.3 | 99.6 | Nasal swab/OP swab/NP swab | 2090 | Yes
Wuxi Biohermes Bio & Medical Technology Co | SARS-CoV-2 Antigen Test Kit (Lateral Flow Assay) | 95.31 | 99.4 | NP swab/OP swab | 2143 | Yes
Xiamen Boson Biotech | Rapid SARS-CoV-2 Antigen Test Card | 93.75 | 99.03 | NP swab | 1278 | Yes
Xiamen Wiz Biotech | SARS-CoV-2 Antigen Rapid Test (Colloidal Gold) | 95.91 | 100 | Nasal swab | 1884 | Yes
Z Biotech Inc. | Anylab Covid-19 Ag Test kit | 95.2 | 100 | NP swab | 2599 | No
Zenum | BENKVO COVID-19 Antigen Rapid Test (Colloidal Gold) kit | 99 | 100 | NP swab | - | No
Zhejiang Anji Saianfu Biotech | reOpenTest COVID-19 Antigen Rapid Test | 98.1/97.5 | 99.1/99.1 | NP swab/Nasal swab | 1295 | Yes
Zhejiang Orient Gene Biotech | Coronavirus Ag Rapid Test Cassette (Swab) | 96.3 | 99.1 | NP swab | 1343 | Yes
Zhejiang Orient Gene Biotech (distributed by Bio-Rad and Menarini) | Coronavirus Ag Rapid Test Cassette (Swab) | 98.3/97.3 | 99.6/100 | NP swab/Nasal swab | 1343 | Yes
Zhuhai Lituo Biotechnology | COVID-19 Antigen Detection Kit (Colloidal Gold) | 96.4/96.9/96.0/93.3 | 99.5 | OP swab/NP swab/Nasal swab/Saliva | 1957 | Yes
Abbreviations
CLIA | Chemiluminescent Immunoassay
CMIA | Chemiluminescent Microparticle Immunoassay
ECLIA | Electrochemiluminescence Immunoassay
ELFA | Enzyme-linked Fluorescent Assay
ELISA | Enzyme-linked Immunosorbent Assay
NA | Not applicable
NCP | Nucleocapsid Protein
NP | Nasopharyngeal
OP | Oropharyngeal
RBD | Receptor-binding Domain
S | Spike Protein
S1 & S2 | S1 & S2 subunit